ab102507 Zinc Quantification Kit

Instructions for Use

For the rapid, sensitive and accurate measurement of Zinc levels in various samples

This product is for research use only and is not intended for diagnostic use.

PLEASE NOTE: With the acquisition of BioVision by Abcam, we have made some changes to component names and packaging to better align with our global standards as we work towards environmental-friendly and efficient growth. You are receiving the same high-quality products as always, with no changes to specifications or protocols.

# Table of Contents

Table of Contents	2

1.	Overview	2

2.	Protocol Summary	4

3.	Materials Supplied	5

4.	Storage and Stability	5

5.	Materials Required, Not Supplied	6

6.	Assay Protocol	7

7.	Data Analysis	9

8.	Troubleshooting	11

# Overview

Zinc, a metallic chemical element, symbol Zn and atomic number 30 is chemically similar to Magnesium due to its similar size and sole oxidation state of 2+. Zinc is an essential mineral of great biological significance, because many enzymes require it as an essential cofactor.

Examples of zinc’s biological roles include signal transduction, gene expression, and regulation of apoptosis, synaptic plasticity and prostate gland function.

Abcam’s Zinc Quantification Kit is a convenient colorimetric assay in which Zinc binds to a ligand that will developed and detected at an absorbance at OD560 nm. The assay can be used with biological samples such as serum, plasma, CSF or urine with detection sensitivity 0.2 µg/ml (~1 – 3 µM).

# Protocol Summary

Sample Preparation

Standard Curve Preparation

Prepare and Add Reaction Mix

Measure Optical Density

# Materials Supplied

| Item | Quantity |
| --- | --- |
| Zinc Reagent 1 | 16 mL |
| Zinc Reagent 2 | 4 mL |
| TCA (7%)/7% TCA | 5 mL |
| Zinc Standard/Zinc Standard (50 mM) | 100 µL |

# Storage and Stability

Upon arrival, store the kit at +4°C and protected from light.

Read the entire protocol before performing the assay.

The reagents are ready to use as supplied.

- Synthetic rubber and glass can contain zinc which may leach into samples. For highest accuracy all glassware should be washed with dilute HCl, rinsed with distilled water and dried prior to use.

- Sample tubes such as Vacutainer® and similar devices should be sealed with Parafilm® rather than the butyl rubber stopper.

- Chelators such as EDTA will give artificially low zinc levels and should be avoided in lysis buffers for all samples. For blood samples, heparin, citrate and oxalate are acceptable anticoagulants.

- Most blood zinc (80%) is contained in erythrocytes and hemolysis will release large amounts into the serum. Abnormally high serum values obtained suggest the collection of another sample and re-testing.

# Materials Required, Not Supplied

- Distilled water or MilliQ

- Microcentrifuge

- Pipettes and pipette tips

- Colorimetric microplate reader

- 96 well plate

- Orbital shaker

- (Optional): HCl, for urine samples

# Assay Protocol

- Sample Preparation:

- Cell lysates and tissue extracts: harvest cells or homogenize tissue with an EDTA-free lysis buffer. Spin down to get rid of cell debris and transfer supernatant to clean tubes.

Lysates generally contain significant amounts of protein, so they should be deproteinized by adding 50 µl of the TCA (7%)/7% TCA solution per 50 µl of the sample (add 1x volume TCA to 1x volume sample).

Spin lysates-TCA mix at top speed for 5 minutes.

Add 20 – 50 µl of sample to a 96-well plate; if volume of cell is < 50 µl, bring the volume to 50 µl/well with dH2O.

- Urine samples: urine samples should be acidified to pH 3 – 4 to dissolve any sediment that could bind to zinc. Add 1 – 2 drops of concentrated HCl per 15 ml sample.

Acidified urine can be used directly in the assay (20 – 50 µl sample/well).

- Other liquid samples (Cell culture media, plasma, serum and other biological fluids): liquid samples can be assayed directly. However, samples that contain significant amounts of protein such as serum, plasma and CSF should be deproteinized prior the assay.

Add 1x volume of the TCA (7%)/7% TCA solution to 1x volume of sample (for example, 50 µl of TCA (7%)/7% TCA to 50 µl of sample).

Spin sample-TCA mix at at top speed for 5 minutes.

Add 20-50 μl of the sample(s) to a 96 well plate; if volume of cells is < 50 µl, bring the volume to 50 μl/well with dH2O.

For unknown samples, we suggest including several dilutions for each sample so that the reading will be within the standard curve range.

Some samples might need the neutralization step following the addition of TCA solution.

- Standard Curve Preparation:

Prepare a 0.5mM Zinc Standard by adding 10 μl of the 50 mM Zinc Standard to 990 μl of dH2O and mixing well.

Add 0, 2, 4, 6, 8, 10 μl into a series of wells. Adjust volume to 50 μl/well with dH2O to generate 0, 1, 2, 3, 4, 5 nmol/well of the Zinc Standard.

- Reaction Mix:

Add 4x parts of Zinc Reagent 1 to 1x part Zinc Reagent 2. Make only as much Zinc reaction mix as is needed for samples and standards to be run. Each sample or standard requires 200 μl of reagent mix. Once mixed, the Zinc reaction mix is good for 2 days at room temperature or 1 week at +4°C.

Add 200 μl of Zinc reaction mix to each standard and sample; incubate 10 min at room temperature.

- Measurement:

Measure assay at OD560nm in a microplate reader.

# Data Analysis

Correct background by subtracting the value derived from the zero Zinc Standard from all readings. The background reading can be significant and must be subtracted.

Plot the Zinc Standard curve.

Read Zinc sample concentrations from the standard curve:

Concentration = Sa / Sv nmol/μl or mM,

Where

Sa is the sample amount (in nmol) from standard curve

Sv is the sample volume (μl) added into the wells

Zinc MW: 65.384 g/mol

# Troubleshooting

| Problem | Reason | Solution |
| --- | --- | --- |
| Assay not working | Assay buffer at wrong temperature | Assay buffer must not be chilled - needs to be at RT |
|  | Protocol step missed | Re-read and follow the protocol exactly |
|  | Plate read at incorrect wavelength | Ensure you are using appropriate reader and filter settings (refer to datasheet) |
|  | Unsuitable microtiter plate for assay | Fluorescence: Black plates (clear bottoms); Luminescence: White plates; Colorimetry: Clear plates. If critical, datasheet will indicate whether to use flat- or U-shaped wells |
| Unexpected results | Measured at wrong wavelength | Use appropriate reader and filter settings described in datasheet |
|  | Samples contain impeding substances | Troubleshoot and also consider deproteinizing samples |
|  | Unsuitable sample type | Use recommended samples types as listed on the datasheet |
|  | Sample readings are outside linear range | Concentrate/ dilute samples to be in linear range |

| Samples with inconsistent readings | Unsuitable sample type | Refer to datasheet for details about incompatible samples |
| --- | --- | --- |
|  | Samples prepared in the wrong buffer | Use the assay buffer provided (or refer to datasheet for instructions) |
|  | Samples not deproteinized (if indicated on datasheet) | Use the 10kDa spin column (ab93349) |
|  | Cell/ tissue samples not sufficiently homogenized | Increase sonication time/ number of strokes with the Dounce homogenizer |
|  | Too many freeze-thaw cycles | Aliquot samples to reduce the number of freeze-thaw cycles |
|  | Samples contain impeding substances | Troubleshoot and also consider deproteinizing samples |
|  | Samples are too old or incorrectly stored | Use freshly made samples and store at recommended temperature until use |
| Lower/ Higher readings in samples and standards | Not fully thawed kit components | Wait for components to thaw completely and gently mix prior use |
|  | Out-of-date kit or incorrectly stored reagents | Always check expiry date and store kit components as recommended on the datasheet |
|  | Reagents sitting for extended periods on ice | Try to prepare a fresh reaction mix prior to each use |
|  | Incorrect incubation time/ temperature | Refer to datasheet for recommended incubation time and/ or temperature |
|  | Incorrect amounts used | Check pipette is calibrated correctly (always use smallest volume pipette that can pipette entire volume) |
| Problem | Reason | Solution |
| Standard curve is not linear | Not fully thawed kit components | Wait for components to thaw completely and gently mix prior use |
|  | Pipetting errors when setting up the standard curve | Try not to pipette too small volumes |
|  | Incorrect pipetting when preparing the reaction mix | Always prepare a master mix |
|  | Air bubbles in wells | Air bubbles will interfere with readings; try to avoid producing air bubbles and always remove bubbles prior to reading plates |
|  | Concentration of standard stock incorrect | Recheck datasheet for recommended concentrations of standard stocks |
|  | Errors in standard curve calculations | Refer to datasheet and re-check the calculations |
|  | Use of other reagents than those provided with the kit | Use fresh components from the same kit |

For further technical questions please do not hesitate to contact us by email (technical@abcam.com) or phone (select “contact us” on www.abcam.com for the phone number for your region).

Technical Support

Copyright © 2023 Abcam. All Rights Reserved. The Abcam logo is a registered trademark. All information / detail is correct at time of going to print.

For all technical or commercial enquiries please go to:

www.abcam.com/contactus

www.abcam.cn/contactus (China)

www.abcam.co.jp/contactus (Japan)