ab108803 – Human Apolipoprotein AI ELISA Kit (APOA1)

Instructions for Use

For the quantitative measurement of Human Apolipoprotein AI (APOA1) in plasma, serum, urine, saliva, milk, CSF, cell culture, cell lysate, and tissue samples.

This product is for research use only and is not intended for diagnostic use.

Table of Contents

INTRODUCTION

1.	BACKGROUND	2

2.	ASSAY SUMMARY	4

3.	PRECAUTIONS	5

GENERAL INFORMATION

4.	STORAGE AND STABILITY	5

5.	MATERIALS SUPPLIED	5

6.	MATERIALS REQUIRED, NOT SUPPLIED	6

7.	LIMITATIONS	6

8.	TECHNICAL HINTS	7

ASSAY PREPARATION

9.	REAGENT PREPARATION	8

10.	STANDARD PREPARATIONS	11

11.	SAMPLE PREPARATION	14

12.	PLATE PREPARATION	16

ASSAY PROCEDURE

13.	ASSAY PROCEDURE	17

DATA ANALYSIS

14.	CALCULATIONS	19

15.	TYPICAL DATA	19

16.	TYPICAL SAMPLE VALUES	20

17.	ASSAY SPECIFICITY	21

RESOURCES

18.	TROUBLESHOOTING	22

19.	NOTES	24

## BACKGROUND

Abcam’s Human Apolipoprotein AI ELISA Kit (APOA1) is designed for the quantitative measurement of apolipoprotein AI concentrations in plasma, serum, urine, saliva, milk, CSF, tissue, cell lysate, and cell culture samples.

An Apolipoprotein AI specific antibody has been precoated onto 96-well plates and blocked. Standards or test samples are added to the wells and subsequently an Apolipoprotein AI specific biotinylated detection antibody is added and then followed by washing with wash buffer. Streptavidin-Peroxidase Complex is added and unbound conjugates are washed away with wash buffer. TMB is then used to visualize Streptavidin-Peroxidase enzymatic reaction. TMB is catalyzed by Streptavidin-Peroxidase to produce a blue color product that changes into yellow after adding acidic stop solution. The density of yellow coloration is directly proportional to the amount of Apolipoprotein AI captured in plate.

Human apolipoprotein AI comprises about 70% of the high-density lipoproteins (HDL) protein mass and apolipoprotein AII another 15–20%. APOA1, a 243-amino acid molecule that contains a series of highly homologous amphipathic alpha-helices, is a 28-kDa single polypeptide that lacks glycosylation or disulfide linkages. About 5–10% of Human plasma apolipoprotein AI exists in a lipoprotein-unassociated state. Apolipoprotein AI appears to have effects on reverse cholesterol transport, and anti-inflammation. Oxidation of specific amino acid residues in apolipoprotein AI may impair cholesterol efflux from macrophages. A majority of HDL functionality is derived from the ability of apolipoprotein AI to sequester phospholipids and cholesterol and interact with plasma enzymes and cellular receptors. During reverse cholesterol transport, HDL interacts with lecithin: cholesteryl acyltransferase (LCAT) and cellular receptors, including ATP-binding cassette transporter protein AI (ABCA1) and the scavenger receptor class B type I in an ordered fashion that is reflected by HDL particle lipid composition. A high-affinity HDL receptor for apolipoprotein AI is beta-chain of ATP synthase on the surface of hepatocytes. The plasma concentration of apolipoprotein AI is one of the best indicators of susceptibility to cardiovascular disease.

## ASSAY SUMMARY

Prepare all reagents, samples and standards as instructed.

Add standard or sample to each well used. Incubate at room temperature.

Wash and add prepared biotin antibody to each well. Incubate at room temperature.

Wash and add prepared Streptavidin-Peroxidase Conjugate. Incubate at room temperature.

Add Chromogen Substrate to each well. Incubate at room temperature. Add Stop Solution to each well. Read immediately.

## PRECAUTIONS

Please read these instructions carefully prior to beginning the assay.

Modifications to the kit components or procedures may result in loss of performance.

## STORAGE AND STABILITY

Store kit at 4°C immediately upon receipt, apart from the SP Conjugate & Biotinylated Antibody, which should be stored at -20°C.

Refer to list of materials supplied for storage conditions of individual components. Observe the storage conditions for individual prepared components in sections 9 & 10.

## MATERIALS SUPPLIED

| Item | Amount | Storage Condition (Before Preparation) |
| --- | --- | --- |
| Apolipoprotein AI Microplate (12 x 8 well strips) | 96 wells | 4°C |
| Apolipoprotein AI Standard | 1 vial | -20°C |
| 10X Diluent N Concentrate | 30 mL | 4°C |
| Biotinylated Human Apolipoprotein AI Antibody | 1 vial | -20°C |
| 100X Streptavidin-Peroxidase Conjugate (SP Conjugate) | 80 µL | -20°C |
| Chromogen Substrate | 7 mL | 4°C |
| Stop Solution | 11 mL | 4°C |
| 20X Wash Buffer Concentrate | 2 x 30 mL | 4°C |
| Sealing Tapes | 3 | N/A |

## MATERIALS REQUIRED, NOT SUPPLIED

These materials are not included in the kit, but will be required to successfully utilize this assay:

- 1 Microplate reader capable of measuring absorbance at 450 nm.

- Precision pipettes to deliver 1 µL to 1 mL volumes.

- Adjustable 1-25 mL pipettes for reagent preparation.

- 100 mL and 1 liter graduated cylinders.

- Absorbent paper.

- Distilled or deionized water.

- Log-log graph paper or computer and software for ELISA data analysis.

- 9 tubes to prepare standard or sample dilutions.

## LIMITATIONS

- Do not mix or substitute reagents or materials from other kit lots or vendors.

## TECHNICAL HINTS

- Samples generating values higher than the highest standard should be further diluted in the appropriate sample dilution buffers.

- Avoid foaming or bubbles when mixing or reconstituting components.

- Avoid cross contamination of samples or reagents by changing tips between sample, standard and reagent additions.

- Ensure plates are properly sealed or covered during incubation steps.

- Complete removal of all solutions and buffers during wash steps.

- This kit is sold based on number of tests. A ‘test’ simply refers to a single assay well. The number of wells that contain sample, control or standard will vary by product. Review the protocol completely to confirm this kit meets your requirements. Please contact our Technical Support staff with any questions.

## REAGENT PREPARATION

Equilibrate all reagents to room temperature (18-25°C) prior to use. Prepare fresh reagents immediately prior to use. When diluting the concentrate, make sure to rinse the bottle thoroughly to extract any precipitates left in the bottle.  Mix the 1x solution gently until the crystals have completely dissolved.

  - 1X Diluent N

Dilute the 10X Diluent N Concentrate 1:10 with reagent grade water. Mix gently and thoroughly. Store for up to 1 month at 4°C.

  - 1X Wash Buffer

Dilute the 20X Wash Buffer Concentrate 1:20 with reagent grade water. Mix gently and thoroughly.

  - 1X Biotinylated Apolipoprotein AI Detector Antibody

    - The stock Biotinylated Apolipoprotein AI Antibody must be diluted with 1X Diluent N according to the label concentration to prepare 1X Biotinylated Apolipoprotein AI Antibody for use in the assay procedure. Observe the label for the “X” concentration on the vial of Biotinylated Apolipoprotein AI Antibody.

    - Calculate the necessary amount of 1X Diluent N to dilute the Biotinylated Apolipoprotein AI Antibody to prepare a 1X Biotinylated Apolipoprotein AI Antibody solution for use in the assay procedure according to how many wells you wish to use and the following calculation:

| Number of Wells Strips | Number of Wells | (VT) Total Volume of 1X Biotinylated Antibody (µL) |
| --- | --- | --- |
| 4 | 32 | 1,760 |
| 6 | 48 | 2,640 |
| 8 | 64 | 3,520 |
| 10 | 80 | 4,400 |
| 12 | 96 | 5,280 |

Any remaining solution should be frozen at -20°C.

Where:

CS = Starting concentration (X) of stock Biotinylated Apolipoprotein AI Antibody (variable)

CF = Final concentration (always = 1X) of 1X Biotinylated Apolipoprotein AI Antibody solution for the assay procedure

VT = Total required volume of 1X Biotinylated Apolipoprotein AI Antibody solution for the assay procedure

VA = Total volume of (X) stock Biotinylated Apolipoprotein AI Antibody

VD = Total volume of 1X Diluent N required to dilute (X) stock Biotinylated Apolipoprotein AI Antibody to prepare 1X Biotinylated Apolipoprotein AI solution for assay procedures

Calculate the volume of (X) stock Biotinylated Antibody required for the given number of desired wells:

(CF / CS) x VT = VA

Calculate the final volume of 1X Diluent N required to prepare the 1X Biotinylated Apolipoprotein AI Antibody:

VT - VA = VD

Example:

NOTE: This example is for demonstration purposes only. Please remember to check your antibody vial for the actual concentration of antibody provided.

CS = 50X Biotinylated Apolipoprotein AI Antibody stock

CF = 1X Biotinylated Apolipoprotein AI Antibody solution for use in the assay procedure

VT = 3,520 µL (8 well strips or 64 wells)

(1X/50X) x 3,520 µL = 70.4 µL

3,520 µL - 70.4 µL = 3,449.6 µL

VA = 70.4 µL total volume of (X) stock Biotinylated Apolipoprotein AI Antibody required

VD = 3,449.6 µL total volume of 1X Diluent N required to dilute the 50X stock Biotinylated Antibody to prepare 1X Biotinylated Apolipoprotein AI Antibody solution for assay procedures

    - First spin the Biotinylated Apolipoprotein AI Antibody vial to collect the contents at the bottom.

    - Add calculated amount VA of stock Biotinylated Apolipoprotein AI Antibody to the calculated amount VD of 1X Diluent N. Mix gently and thoroughly.

  - 1X SP Conjugate

Spin down the 100X Streptavidin-Peroxidase Conjugate (SP Conjugate) briefly and dilute the desired amount of the conjugate 1:100 with 1X Diluent N.

Any remaining solution should be frozen at -20°C.

## STANDARD PREPARATIONS

- Prepare serially diluted standards immediately prior to use. Always prepare a fresh set of standards for every use.

- Any remaining standard should be stored at -20°C after reconstitution and used within 5 days.

- This procedure prepares sufficient standard dilutions for duplicate wells.

- Reconstitution of the Apolipoprotein AI Standard vial to prepare a 200 ng/mL Apolipoprotein AI Stock Standard:

    - First consult the Apolipoprotein AI Standard vial to determine the mass of protein in the vial.

    - Calculate the appropriate volume of 1X Diluent N to add when resuspending the Apolipoprotein AI Standard vial to produce a 200 ng/mL Apolipoprotein AI Stock Standard by using the following equation:

CS = Starting mass of Apolipoprotein AI Standard (see vial label) (µg)

CF = 200 ng/mL Apolipoprotein AI Stock Standard final required concentration

VD = Required volume of 1X Diluent N for reconstitution (µL)

Calculate total required volume 1X Diluent N for resuspension:

(CS/ CF) x 1,000 = VD

Example:

NOTE: This example is for demonstration purposes only. Please remember to check your standard vial for the actual amount of standard provided.

CS = 180 ng of Apolipoprotein AI Standard in vial

CF = 200 ng/mL Apolipoprotein AI Stock Standard final concentration

VD = Required volume of 1X Diluent N for reconstitution

(180 ng / 200 ng/mL) x 1,000 = 900 µL

    - First briefly spin the Apolipoprotein AI Standard Vial to collect the contents on the bottom of the tube.

    - Reconstitute the Apolipoprotein AI Standard vial by adding the appropriate calculated amount VD of 1X Diluent N to the vial to generate the 200 ng/mL Apolipoprotein AI Stock Standard. Mix gently and thoroughly.

- Allow the reconstituted 200 ng/mL Apolipoprotein AI Stock Standard to sit for 10 minutes with gentle agitation prior to making subsequent dilutions

- Label eight tubes #2-8.

- Prepare the 100 ng/mL Standard #2 by adding 120 µL of the reconstituted 200 ng/mL Apolipoprotein AI Stock Standard to tube #2.

- Add 120 µL of 1X Diluent N to tube #2 – 8.

- To prepare Standard #3, add 120 μL of the Standard #2 into tube #3 and mix gently.

- To prepare Standard #4, add 120 μL of the Standard #3 into tube #4 and mix gently.

- Using the table below as a guide, prepare subsequent serial dilutions.

- 1X Diluent N serves as the zero standard, 0 ng/mL (tube #8).

Standard Dilution Preparation Table

| Standard # | Volume to Dilute (µL) | Volume Diluent N (µL) | Total Volume (µL) | Starting Conc. (ng/mL) | Final Conc. (ng/mL) |
| --- | --- | --- | --- | --- | --- |
| 1 | Step 10.1 |  |  |  | 200.0 |
| 2 | 120 | 120 | 240 | 200.0 | 100.0 |
| 3 | 120 | 120 | 240 | 100.0 | 50.0 |
| 4 | 120 | 120 | 240 | 50.0 | 25.0 |
| 5 | 120 | 120 | 240 | 25.0 | 12.5 |
| 6 | 120 | 120 | 240 | 12.5 | 6.25 |
| 7 | 120 | 120 | 240 | 6.25 | 3.125 |
| 8 | - | 120 | 120 | - | 0 |

## SAMPLE PREPARATION

  - Plasma

Collect plasma using one tenth volume of 0.1 M sodium citrate as an anticoagulant. Centrifuge samples at 3000 x g for 10 minutes. Dilute samples 100,000-fold into 1X Diluent and assay. Depending on application needs, user should determine proper dilutions. The undiluted samples can be stored at -20°C or below for up to 3 months. Avoid repeated freeze-thaw cycles (EDTA and Heparin can also be used as an anticoagulant).

  - Serum

Samples should be collected into a serum separator tube. After clot formation, centrifuge samples at 3000 x g for 10 minutes and remove serum. Dilute samples 100,000-fold into 1X Diluent N and assay. Depending on application needs, user should determine proper dilutions. The undiluted samples can be stored at -20°C or below for up to 3 months. Avoid repeated freeze-thaw cycles.

  - Urine

Collect urine using sample pot. Centrifuge samples at 800 x g for 10 minutes. Dilute samples at 1:2 into 1X Diluent N or within the range of 1x – 10x, and assay. Depending on application needs, user should determine proper dilutions. The undiluted samples can be stored at -20°C or below for up to 3 months. Avoid repeated freeze-thaw cycles.

  - Cell culture supernatants

Centrifuge cell culture media at 1,500 x g for 10 minutes at 4°C to remove debris and collect supernatants. If necessary, dilute samples into 1X Diluent N; user should determine optimal dilution factor depending on application needs. Store undiluted samples at -20°C or below. Avoid repeated freeze-thaw cycles.

  - Cell Lysate

Rinse cell with cold PBS and then scrape the cell into a tube with 5 ml cold PBS with 0.5 M EDTA. Centrifuge suspension at 1,500 rpm for 10 minutes at 4°C and aspirate supernatant. Re-suspend the pellet in ice-cold Lysis Buffer (PBS, 1% Triton X-100, protease inhibitor cocktail). For every 1 x 10^6 cells add approximately 100 μL of ice-cold Lysis Buffer. Incubate on ice for 60 min. Centrifuge at 13,000 rpm for 30 minutes at 4°C and collect supernatant for assay. If necessary, dilute samples into 1X Diluent N; user should determine optimal dilution factor depending on application needs.

  - Saliva

Collect saliva using sample tube. Centrifuge samples at 
800 x g for 10 minutes. Dilute samples at 1:32 into 1X Diluent N, or within the range of 3x – 300x, and assay. The undiluted samples can be stored at -20°C or below for up to 3 months. Avoid repeated freeze-thaw cycles.

  - CSF

Collect cerebrospinal fluid (CSF) using sample pot. Centrifuge samples at 3000 x g for 10 minutes. Dilute samples at 1:400 into 1X Diluent N or within the range of 5x – 500x and assay. The undiluted samples can be stored at -80°C for up to 3 months. Avoid repeated freeze-thaw cycles.

  - Milk

Collect milk using sample tube. Centrifuge samples at 
800 x g for 10 minutes. Use undiluted samples or dilute into 1X Diluent N 320x or within the range of 30x – 3000x, and assay. Store undiluted samples at -20°C or below for up to 3 months. Avoid repeated freeze-thaw cycles.

  - Tissue:

Extract tissue samples with 0.1 M phosphate-buffered saline (pH 7.4) containing 1% Triton X-100 and centrifuge at 14000 x g for 20 minutes.  Collect the supernatant and measure the protein concentration. If necessary, dilute samples into MIX Diluent; user should determine optimal dilution factor depending on application needs.  Store remaining extract at -80°C.  Avoid repeated freeze-thaw cycles.

## PLATE PREPARATION

- The 96 well plate strips included with this kit are supplied ready to use. It is not necessary to rinse the plate prior to adding reagents.

- Unused well plate strips should be returned to the plate packet and stored at 4°C.

- For statistical reasons, we recommend each sample should be assayed with a minimum of two replicates (duplicates).

- Well effects have not been observed with this assay. Contents of each well can be recorded on the template sheet included in the Resources section.

## ASSAY PROCEDURE

●	Equilibrate all materials and prepared reagents to room temperature (18 - 25°C) prior to use.

●	It is recommended to assay all standards, controls and samples in duplicate.

  - Prepare all reagents, working standards and samples as instructed. Equilibrate reagents to room temperature before use. The assay is performed at room temperature (18-25°C).

  - Remove excess microplate strips from the plate frame and return them immediately to the foil pouch with desiccant inside. Reseal the pouch securely to minimize exposure to water vapor and store in a vacuum desiccator.

  - Add 50 μL of Apolipoprotein AI Standard or sample per well. Cover wells with a sealing tape and incubate for two hours. Start the timer after the last sample addition.

  - Wash five times with 200 μL of 1X Wash Buffer manually. Invert the plate each time and decant the contents; tap it 4-5 times on absorbent paper towel to completely remove the liquid. If using a machine, wash six times with 300 μL of 1X Wash Buffer and then invert the plate, decant the contents; tap it 4-5 times on absorbent paper towel to completely remove the liquid.

  - Add 50 μL of 1X Biotinylated Apolipoprotein AI Antibody to each well and incubate for one hour.

  - Wash microplate as described above.

  - Add 50 μL of 1X SP Conjugate to each well and incubate for 30 minutes. Turn on the microplate reader and set up the program in advance.

  - Wash microplate as described above.

  - Add 50 μL of Chromogen Substrate per well and incubate in ambient light for about 20 minutes or till the optimal blue colour density develops. Gently tap plate to ensure thorough mixing and break the bubbles in the well with pipette tip.

  - Add 50 μL of Stop Solution to each well. The color will change from blue to yellow.

  - Read the absorbance on a microplate reader at a wavelength of 450 nm immediately. If wavelength correction is available, subtract readings at 570 nm from those at 450 nm to correct optical imperfections. Otherwise, read the plate at 450 nm only. Please note that some unstable black particles may be generated at high concentration points after stopping the reaction for about 10 minutes, which will reduce the readings.

## CALCULATIONS

Calculate the mean value of the triplicate readings for each standard and sample. To generate a Standard Curve, plot the graph using the standard concentrations on the x-axis and the corresponding mean 450 nm absorbance on the y-axis. The best-fit line can be determined by regression analysis using log-log or four-parameter logistic curve-fit. Determine the unknown sample concentration from the Standard Curve and multiply the value by the dilution factor.

## TYPICAL DATA

TYPICAL STANDARD CURVE – Data provided for demonstration purposes only. A new standard curve must be generated for each assay performed.

## TYPICAL SAMPLE VALUES

SENSITIVITY –

The minimum detectable dose of Apolipoprotein AI is typically 1 ng/mL.

RECOVERY –

Standard Added Value: 12.5 – 100 ng/mL

Recovery %: 91 – 108.

Average Recovery %: 98

LINEARITY OF DILUTION –

|  | Average % Expected Value |  |
| --- | --- | --- |
| Sample Dilution | Plasma | Serum |
| 25000X | 109 | 107 |
| 50000X | 95 | 97 |
| 100000X | 95 | 93 |

PRECISION –

|  | Intra- Assay | Inter- Assay |
| --- | --- | --- |
| % CV | 5.6 | 10.5 |

## ASSAY SPECIFICITY

| Species | % Cross Reactivity |
| --- | --- |
| Canine | None |
| Bovine | None |
| Equine | None |
| Monkey | <5 |
| Mouse | None |
| Rat | None |
| Swine | None |
| Rabbit | None |

| Protein | % Cross Reactivity |
| --- | --- |
| Apo-CIII | <10 |

No significant cross reactivity observed with Apo-AII, Apo-AIV, Apo-AV, Apo-B, Apo-CII, Apo-E, and Apo-M.

10% FBS in culture media will not affect the assay.

## TROUBLESHOOTING

| Problem | Cause | Solution |
| --- | --- | --- |
| Poor standard curve | Improper standard dilution | Confirm dilutions made correctly |
|  | Standard improperly reconstituted (if applicable) | Briefly spin vial before opening; thoroughly resuspend powder (if applicable) |
|  | Standard degraded | Store sample as recommended |
|  | Curve doesn't fit scale | Try plotting using different scale |
| Low signal | Incubation time too short | Try overnight incubation at 4oC |
|  | Target present below detection limits of assay | Decrease dilution factor; concentrate samples |
|  | Precipitate can form in wells upon substrate addition when concentration of target is too high | Increase dilution factor of sample |
|  | Using incompatible sample type (e.g. serum vs. cell extract) | Detection may be reduced or absent in untested sample types |
|  | Sample prepared incorrectly | Ensure proper sample preparation/dilution |
| Large CV | Bubbles in wells | Ensure no bubbles present prior to reading plate |
|  | All wells not washed equally/thoroughly | Check that all ports of plate washer are unobstructed wash wells as recommended |
|  | Incomplete reagent mixing | Ensure all reagents/master mixes are mixed thoroughly |
|  | Inconsistent pipetting | Use calibrated pipettes and ensure accurate pipetting |
|  | Inconsistent sample preparation or storage | Ensure consistent sample preparation and optimal sample storage conditions (eg. minimize freeze/thaws cycles) |

| Problem | Cause | Solution |
| --- | --- | --- |
| High background/ Low sensitivity | Wells are insufficiently washed | Wash wells as per protocol recommendations |
|  | Contaminated wash buffer | Make fresh wash buffer |
|  | Waiting too long to read plate after adding STOP solution | Read plate immediately after adding STOP solution |
|  | Improper storage of ELISA kit | Store all reagents as recommended. Please note all reagents may not have identical storage requirements. |
|  | Using incompatible sample type (e.g. Serum vs. cell extract) | Detection may be reduced or absent in untested sample types |

## NOTES

Technical Support

Copyright © 2023 Abcam. All Rights Reserved. The Abcam logo is a registered trademark. All information / detail is correct at time of going to print.

For all technical or commercial enquiries please go to:

www.abcam.com/contactus

www.abcam.cn/contactus (China)

www.abcam.co.jp/contactus (Japan)