ab115091 – Histone H3 Total Quantification Kit (Colorimetric)

Instructions for Use

For the measurement of Histone H3 proteins independent of its modified state

This product is for research use only and is not intended for diagnostic use.

Table of Contents

INTRODUCTION

1.	BACKGROUND	2

2.	ASSAY SUMMARY	4

GENERAL INFORMATION

3.	PRECAUTIONS	5

4.	STORAGE AND STABILITY	5

5.	MATERIALS SUPPLIED	6

6.	MATERIALS REQUIRED, NOT SUPPLIED	6

7.	LIMITATIONS	7

8.	TECHNICAL HINTS	7

ASSAY PREPARATION

9.	REAGENT PREPARATION	8

10.	SAMPLE PREPARATION	8

ASSAY PROCEDURE

11.	ASSAY PROCEDURE	10

DATA ANALYSIS

12.	ANALYSIS	11

RESOURCES

13.	TROUBLESHOOTING	12

14.	NOTES	13

## BACKGROUND

Histone H3, along with H2A, H2B, and H4, is involved in the structure of chromatin in eukaryotic cells. Histone H3 can undergo several different types of epigenetic modifications that influence cellular processes such as transcription activation/ inactivation, chromosome packaging, and DNA damage/repair. These modifications including acetylation, phosphorylation, methylation, ubiquitination, and ADP-ribosylation, occur on the N-terminal tail domains of histone H3 through catalyzing of histone modifying enzymes, which result in remodeling of the nucleosome structure into an open conformation more accessible to transcription complexes. In most species, histone H3 is primarily acetylated at lysine 9, 14, 18, 23, and 56, and methylated at lysine 4, 9, 27, 36 and 79, and phosphorylated at ser10, ser28, Thr3, and Thr11, respectively. Thus, quantitative detection of various histone modifications would provide useful information for better understanding epigenetic regulation of cellular processes and for developing HMT-targeted drugs.

Abcam provides a series of kits used for quantifying all sites/degrees of histone H3 modification. For added convenience and more quantitative interpretation of results, we provide the Histone H3 Total Quantification Kit (Colorimetric).

ab115091 is designed for quantifying levels of histone H3 proteins independent of its modified state and can also be used for normalizing the modified histone H3 content of samples when run in parallel with Abcam histone modification quantification kit series.

This kit has the following features:

- Quick and efficient procedure, which can be finished within 2.5 hours

- Innovative colorimetric assay without the need for radioactivity, electrophoresis, or chromatography

- Specifically captures histone H3 with the detection limit as low as 10 ng/well and detection range from 20 ng-1 μg/well of histone extracts

- The unmodified histone H3 control is conveniently included for the quantification of the amount of total histone H3

- Strip microplate format makes the assay flexible: manual or high throughput

- Simple, reliable, and consistent assay conditions

Abcam’s Histone H3 Total Quantification Kit (Colorimetric) is designed for quantifying levels of histone H3 proteins independent of its modified state. In an assay with this kit, the histone H3 protein is captured to the strip wells coated with an anti-histone H3 antibody. The captured histone H3 can then be detected with a detection antibody, followed by a color development reagent. The ratio of histone H3 is proportional to the intensity of absorbance. The absolute amount histone H3 can be quantified by comparing to the standard control.

## ASSAY SUMMARY

Tissue disaggregation or cell lysis

Histone extracts

Histone H3 bound to assay wells cross link

Add detection antibody after wash

Add color developing solution for color development and measure absorbance

## PRECAUTIONS

Please read these instructions carefully prior to beginning the assay.

All kit components have been formulated and quality control tested to function successfully as a kit. Modifications to the kit components or procedures may result in loss of performance.

## STORAGE AND STABILITY

Store kit as given in the table and away from light upon receipt.

Observe the storage conditions for individual prepared components in sections 9 & 10.

For maximum recovery of the products, centrifuge the original vial prior to opening the cap.

Check if 10X Wash Buffer and Antibody Buffer contain salt precipitates before use. If so, warm at room temperature or 37°C and shake the buffer until the salts are re-dissolved.

## MATERIALS SUPPLIED

| Item | 48 Tests | 96 Tests | Storage Condition (Before Preparation) |
| --- | --- | --- | --- |
| 10X Wash Buffer | 10 mL | 20 mL | 4°C |
| Antibody Buffer | 6 mL | 12 mL | 4°C |
| Detection Antibody, 1 mg/mL | 5 µL | 10 µL | -20°C |
| Color Developer | 5 mL | 10 mL | 4°C |
| Stop Solution | 3 mL | 6 mL | 4°C |
| Standard Control, 100 μg/mL | 10 µL | 20 µL | -20°C |
| 8-Well Assay Strip (with Frame) | 4 | 9 | 4°C |
| 8-Well Standard Control Strips* | 2 | 3 | 4°C |

*These wells are identified by a green ring around the top.

## MATERIALS REQUIRED, NOT SUPPLIED

These materials are not included in the kit, but will be required to successfully utilize this assay:

- Pipettes and pipette tips

- Orbital shaker

- Reagent reservoir

- Microplate reader

## LIMITATIONS

- Assay kit intended for research use only. Not for use in diagnostic procedures

- Do not use kit or components if it has exceeded the expiration date on the kit labels

- Do not mix or substitute reagents or materials from other kit lots or vendors. Kits are QC tested as a set of components and performance cannot be guaranteed if utilized separately or substituted

- Any variation in operator, pipetting technique, washing technique, incubation time or temperature, and kit age can cause variation in binding

## TECHNICAL HINTS

- Avoid foaming or bubbles when mixing or reconstituting components

- Avoid cross contamination of samples or reagents by changing tips between sample, standard and reagent additions

- Ensure plates are properly sealed or covered during incubation steps

- Complete removal of all solutions and buffers during wash steps

- This kit is sold based on number of tests.  A ‘test’ simply refers to a single assay well. The number of wells that contain sample, control or standard will vary by product. Review the protocol completely to confirm this kit meets your requirements. Please contact our Technical Support staff with any questions

## REAGENT PREPARATION

All reagents provided are ready to use.

## SAMPLE PREPARATION

Prepare histone extracts from cells/tissues treated or untreated by using your own successful method (acid extraction or high salt extraction). For your convenience and the best results, Abcam offers the Histone Extraction Kit (ab113476) optimized for use in Abcam’s modified histone quantification series. Alternatively, preparation of histone extracts can also be performed using the procedure below:

  - For tissues (treated and untreated), weigh the sample and cut the sample into small pieces (1-2 mm3) with a scalpel or scissors. Transfer tissue pieces to a Dounce homogenizer. Add TEB buffer (PBS containing 0.5% Triton X 100, 2 mM PMSF and 0.02% NaN3) at 200 mg/ml, and disaggregate tissue pieces by 50-60 strokes. Transfer homogenized mixture to a 15 mL conical tube and centrifuge at 3000 rpm for 5 minutes at 4°C. If total mixture volume is less than 2 mL, transfer mixture to a 2 mL vial and centrifuge at 10,000 rpm for 1 minute at 4°C. Remove supernatant.

For cells (treated and untreated), harvest cells and pellet the cells by centrifugation at 1000 rpm for 5 minutes at 4°C. Resuspend cells in TEB buffer at 107 cells/mL and lyse cells on ice for 10 minutes with gentle stirring. Centrifuge at 3000 rpm for 5 minutes at 4°C. If total volume is less than 2 mL, transfer cell lysates to a 2 mL vial and centrifuge at 10,000 rpm for 1 minute at 4°C. Remove supernatant.

  - Resuspend cell/tissue pellet in 3 volumes (approx. 200 μL/107 cells or 200 mg tissues) of extraction buffer (0.5N HCl + 10% glycerol) and incubate on ice for 30 minutes.

  - Centrifuge at 12,000 rpm for 5 minutes at 4°C and remove the supernatant fraction to a new vial.

  - Add 8 volumes (relative to original pellet size, or approx. 0.6 mL/107 cells or 200 mg tissues) of acetone and leave at -20°C overnight.

  - Centrifuge at 12,000 rpm for 5 minutes and air-dry the pellet. Dissolve the pellet in distilled water (30-50 μL/107 cells or 200 mg tissues).

  - Quantify the protein concentration. Aliquot the extract and store the extract at -20°C or -80°C.

Histone extractions can be used immediately or stored at -80°C for future use.

## ASSAY PROCEDURE

  - Determine the number of strip wells required. Leave these strips in the plate frame (remaining unused strips can be placed back in the bag. Seal the bag tightly and store at 4°C). Dilute the 10X Wash Buffer with distilled water (pH 7.2-7.5) at a 1:10 ratio to make 1X Wash Buffer (e.g. 1 mL of 10X Wash Buffer + 9 mL of water).

  - Add 50 µl of Antibody Buffer into each well. For the sample, add 50-200 ng of the histone extract into the sample wells. For the standard curve, dilute the Standard Control with Antibody Buffer to 1-60 ng/µl at 5-7 points (e.g., 1, 1.9, 3.8, 7.5, 15, 30, and 60 ng/µl). Add 1 µl of Standard Control at the different concentrations into the standard well. For the blank, do not add any nuclear extracts or standard control protein. Mix and cover the strip wells with Parafilm M and incubate at room temperature for 1-2 hours.

  - Aspirate and wash the wells with 150 μL of 1X Wash Buffer three times.

  - Dilute the Detection Buffer (at a 1:1000 ratio) to 1 μg/mL with Antibody Buffer. Add 50 μL of Diluted Detection Antibody to each well and incubate at room temperature for 60 minutes on an orbital shaker (100 rpm).

  - Aspirate and wash the wells with 150 μL of 1X Wash Buffer six times.

  - Add 100 μL of the Color Developer into the wells and incubate at room temperature for 2-10 minutes away from light. Monitor the color development in the sample and standard wells (blue).

  - Add 50 μL of Stop Solution to each well to stop enzyme reaction when the color in the standard wells containing the higher concentrations of standard control turns medium blue. The color should change to yellow and absorbance can be read on a microplate reader at 450 nm within 2-15 minutes.

  - Calculate % histone H3 using the formulae shown in Section 12 – Data Analysis.

## ANALYSIS

Calculate the % H3 using the following formula:

x 100%

Histone H3 % =        Treated (Tested) Sample OD – Blank OD

Untreated (Control) Sample OD – Blank OD

For an accurate calculation, plot OD value versus amount of Standard Control and determine the slope as delta OD/ng.

Calculate the amount of Histone H3 using the following formula:

x 1000

Amount (ng/mg protein) =       Sample OD – Blank OD

Protein (µg)* x Slope

*Histone extract amount added into the sample well at step 11.2.

## TROUBLESHOOTING

| Problem | Cause | Solution |
| --- | --- | --- |
| No Signal for Both the Standard Control and the Samples | Reagents are added incorrectly | Check if reagents are added in order and if some steps of the procedure are omitted by mistake |
|  | Incubation time and temperature is incorrect | Ensure the incubation time and temperature described in the protocol is followed correctly |
| No Signal or Very Weak Signal for Only the Standard Control | The amount of Standard control is not added into the standard control wells or is added insufficiently | Ensure a sufficient amount of control is added to the well |
| No Signal for Only the Sample | The protein sample is not properly extracted | Ensure the procedure and reagents are correct for the nuclear protein extraction |
|  | The protein amount is added into well insufficiently | Ensure extract contains a sufficient amount of protein |
|  | Protein extracts are incorrectly stored | Ensure the protein extracts are stored at -20°C or -80°C |
| High Background Present for the Blank | The well is not washed sufficiently | Check if wash at each step is performed according to the protocol |
|  | Contaminated by the Standard control | Ensure the well is not contaminated from adding the control protein or by using control protein contaminated tips |
|  | Overdevelopment | Decrease development time in step 11.6 |

## NOTES

Technical Support

Copyright © 2024 Abcam. All Rights Reserved. The Abcam logo is a registered trademark. All information / detail is correct at time of going to print.

For all technical or commercial enquiries please go to:

www.abcam.com/contactus

www.abcam.cn/contactus (China)

www.abcam.co.jp/contactus (Japan)