ab174094 Glycerol-3-phosphate (G3P) Assay Kit (Colorimetric)

Instructions for Use

For the sensitive and accurate measurement of G3P activity in a variety of samples

This product is for research use only and is not intended for diagnostic use.

PLEASE NOTE: With the acquisition of BioVision by Abcam, we have made some changes to component names and packaging to better align with our global standards as we work towards environmental-friendly and efficient growth. You are receiving the same high-quality products as always, with no changes to specifications or protocols.

# Table of Contents

Table of Contents	2

1.	Overview	3

2.	Protocol Summary	4

3.	Kits Components	5

4.	Storage and Stability	5

5.	Materials Required, Not Supplied	6

6.	Reagents Preparation	7

7.	Assay Protocol	8

8.	Data Analysis	12

9.	Troubleshooting	15

# Overview

Glycerol-3-phosphate (G3P) is an important intermediate for all living organisms. Glycerol-3-Phosphate is produced either by glycerol via glycerol kinase or by dihydroxyacetone phosphate through glycerol-3-phosphate dehydrogenase. In response to cellular signals, glycerol-3-phosphate can be utilized in multiple pathways: it can be further converted into glyceraldehyde-3-phosphate and enter glycolysis or rapidly generate NAD+ in brain or muscle tissues through the G3P shuttle or enter the lipid biosynthetic pathway. Recent studies have found that glycerol-3-phosphate is a novel regulator and plays a fundamental defense role in plant pathogenesis.

Abcam’s Glycerol-3-Phosphate Assay kit (ab174094) is a sensitive, fast and easy-to-use kit. In this assay, G3P is oxidized by Development Enzyme Mix VII/G3P Enzyme Mix to form an intermediate, which reduces a nearly colorless Developer Solution III/probe to a colored product with strong absorbance at 450 nm. This assay kit can detect G3P less than 2 nmol/well and can be used for a variety of sample types.

Figure 1: Assay Procedure.

# Protocol Summary

Reagent Preparation

Standard Curve Preparation

Control & Samples Preparation

Reaction

Measurement and Calculation

# Kits Components

| Item | Quantity |
| --- | --- |
| Assay Buffer V/G3P Assay Buffer | 25 mL |
| Development Enzyme Mix VII/G3P Enzyme Mix (Lyophilized) | 1 vial |
| Developer Solution III/G3P Probe (Lyophilized) | 1 vial |
| G3P Standard/G3P Standard (Lyophilized) | 1 vial |

# Storage and Stability

Upon arrival, store the kit at -20°C and protect from light. Please read the entire protocol before performing the assay. Avoid repeated freeze/thaw cycles.

Briefly centrifuge all small vials prior to opening.

# Materials Required, Not Supplied

- 96-well clear plate with flat bottom (for colorimetric reading)

- Multi-well spectrophotometer (ELISA reader)

- Multi-channel pipette

- Distilled water

# Reagents Preparation

- Developer Solution III/G3P Probe:

Reconstitute with 220 µL dH2O. Pipette up and down to dissolve completely. Stable for 2 months at -20°C.

- Development Enzyme Mix VII/G3P Enzyme Mix:

Reconstitute with 220 µL Assay Buffer V/G3P Assay Buffer. Pipette up and down to dissolve completely. Keep on ice while in use. Aliquot and store at -20°C. Avoid repeated freeze/thaw.

- G3P Standard:

Reconstitute with 100 µL dH2O to generate a 100 mM (100 nmol/µL) G3P Standard solution. Keep on ice while in use. Store at –20°C. Use within two months.

- Assay Buffer V/G3P Assay Buffer:

Warm Assay Buffer V/G3P Assay Buffer to room temperature before use.

# Assay Protocol

- Control and Sample Preparation:

- Cells (starting material: 1 x 106 cells)

Harvest cells and spin down briefly and discard supernatant. Resuspend the cell pellet in 200 µL ice cold Assay Buffer V/G3P Assay Buffer and put on ice. Homogenize with a Douncer homogenizer (10 – 15 passes) on ice, or by pipetting up and down using a smaller tip, until efficient lysis is confirmed by viewing the cells under the microscope. Centrifuge homogenate at 12000 rpm for 5 minutes at 4°C to remove cell debris and collect the supernatant.

Use the supernatant for your subsequent assays. Test different dilutions of the sample to ensure the readings will fall within the linear range of the standard curve.

Add 1 – 50 µL test sample to the wells of a 96-well plate. If volume need is <50 µL, bring it up to 50 µL with Assay Buffer V/G3P Assay Buffer.

- Tissue Sample (starting material: 10mg tissue)

Cut tissue in small pieces, add 200 µL ice cold Assay Buffer V/G3P Assay Buffer and put on ice. Homogenize using a Douncer homogenizer (10 – 15 passes) on ice, until efficient lysis is confirmed, by viewing the cells under the microscope. Spin down the samples and collect the supernatant.

Use the supernatant for your subsequent assays. Test different dilutions of the sample to ensure the readings will fall within the linear range of the standard curve.

Add 1 – 50 µL test sample to wells of a 96-well plate. If volume needed is <50 µL, bring it up to 50 µL with Assay Buffer V/G3P Assay Buffer.

NOTE:

For unknown samples, we suggest testing several doses of your samples to ensure the readings are within the Standard Curve range.

NADH in samples will generate background. For samples suspected of having high NADH levels, prepare parallel samples well(s) as background control.

- Standard Curve Preparation:

- Prepare a 1mM (1nmol/µL) G3P Standard by adding 10 µL of 100 mM G3P Standard to 990 µL dH2O. Mix well by pipetting up and down.

- Using the 1 mM G3P Standard please prepare a standard curve as follows, in a microplate or microcentrifuge tubes:

| G3P 1 mM amount (µL) | Assay Buffer V/G3P Assay buffer (µL) | Amount in well | End concentration G3P in well |
| --- | --- | --- | --- |
| 0 | 150 | 50 µL | 0 nmol/well |
| 6 | 144 | 50 µL | 2 nmol/well |
| 12 | 138 | 50 µL | 4 nmol/well |
| 18 | 132 | 50 µL | 6 nmol/well |
| 24 | 126 | 50 µL | 8 nmol/well |
| 30 | 120 | 50 µL | 10 nmol/well |

Add 50 µL of each standard dilution into a 96-well plate to set up standard. Each dilution has enough amount of standard to set up 2 x duplicates x 50 µL/well.

- Reaction Mix:

Prepare Reaction Mix for each reaction containing:

|  | Reaction Mix | Background Control Mix |
| --- | --- | --- |
| Assay Buffer V/G3P Assay Buffer | 46 µL | 48 µL |
| Development Enzyme Mix VII/G3P Enzyme Mix | 2 µL | --- |
| Developer Solution III/G3P Probe | 2 µL | 2 µL |

Mix enough reagents for the number of assays to be performed. Prepare a Master Mix of the Reaction Mix to ensure consistency. We recommend the following calculation:

|  | Reaction Mix |
| --- | --- |
| Assay Buffer V/G3P Assay Buffer | 46 µL x (Nb samples + Standards + 1) |
| Development Enzyme Mix VII/G3P Enzyme Mix | 2 µL x (Nb samples + Standards + 1) |
| Developer Solution III/G3P Probe | 2 µL x (Nb samples + Standards + 1) |

We also recommend preparing enough Background Control Mix to set up duplicate readings for the NADH background control using the same calculation.

- Plate set up and Detection:

- Add 50 µL of standard or sample to wells

- Add 50 µL of the Reaction Mix to wells containing standard and samples. Mix well.

- If necessary, also add 50 µL of Background Control Mix to 50 µL of background control well(s). Mix well.

- Incubate for 40 minutes at 37°C.

- Measure OD450nm.

# Data Analysis

Calculations:

  - Correct background by subtracting the value derived from the zero standard from all readings.

  - Plot the G3P Standard Curve.

  - If sample Background Control reading is significant, subtract the Background Control reading from sample reading. Apply the corrected sample reading to G3P Standard Curve to get B (nmol of G3P) in the sample wells.

Sample G3P concentration = B/V x Dilution Factor

= nmol/mL = µM

Where:

B = amount of G3P in the sample well (nmol)

V = volume of sample used in the reaction well (mL)

G3P in samples can also be expressed in nmol/mg of protein.

Glycerol-3-Phosphate molecular weight: 172.074 g/mol

Figure 2: (a) G3P Standard Curve. (b) Measurement of G3P in rat liver (100 µg), 3T3-L1 (40 µg) and Hela (50 µg) lysate. Assays were performed following the kit protocol.

# Troubleshooting

| Problem | Reason | Solution |
| --- | --- | --- |
| Assay not working | Assay buffer at wrong temperature | Assay buffer must not be chilled - needs to be at RT |
|  | Protocol step missed | Re-read and follow the protocol exactly |
|  | Plate read at incorrect wavelength | Ensure you are using appropriate reader and filter settings (refer to datasheet) |
|  | Unsuitable microtiter plate for assay | Fluorescence: Black plates (clear bottoms); Luminescence: White plates; Colorimetry: Clear plates. If critical, datasheet will indicate whether to use flat- or U-shaped wells |
| Unexpected results | Measured at wrong wavelength | Use appropriate reader and filter settings described in datasheet |
|  | Samples contain impeding substances | Troubleshoot and also consider deproteinizing samples |
|  | Unsuitable sample type | Use recommended samples types as listed on the datasheet |
|  | Sample readings are outside linear range | Concentrate/ dilute samples to be in linear range |

| Problem | Reason | Solution |
| --- | --- | --- |
| Samples with inconsistent readings | Unsuitable sample type | Refer to datasheet for details about incompatible samples |
|  | Samples prepared in the wrong buffer | Use the assay buffer provided (or refer to datasheet for instructions) |
|  | Samples not deproteinized (if indicated on datasheet) | Use the 10kDa spin column (ab93349) |
|  | Cell/ tissue samples not sufficiently homogenized | Increase sonication time/ number of strokes with the Dounce homogenizer |
|  | Too many freeze-thaw cycles | Aliquot samples to reduce the number of freeze-thaw cycles |
|  | Samples contain impeding substances | Troubleshoot and also consider deproteinizing samples |
|  | Samples are too old or incorrectly stored | Use freshly made samples and store at recommended temperature until use |
| Lower/ Higher readings in samples and standards | Not fully thawed kit components | Wait for components to thaw completely and gently mix prior use |
|  | Out-of-date kit or incorrectly stored reagents | Always check expiry date and store kit components as recommended on the datasheet |
|  | Reagents sitting for extended periods on ice | Try to prepare a fresh reaction mix prior to each use |
|  | Incorrect incubation time/ temperature | Refer to datasheet for recommended incubation time and/ or temperature |
|  | Incorrect amounts used | Check pipette is calibrated correctly (always use smallest volume pipette that can pipette entire volume) |
| Problem | Reason | Solution |
| Standard curve is not linear | Not fully thawed kit components | Wait for components to thaw completely and gently mix prior use |
|  | Pipetting errors when setting up the standard curve | Try not to pipette too small volumes |
|  | Incorrect pipetting when preparing the reaction mix | Always prepare a master mix |
|  | Air bubbles in wells | Air bubbles will interfere with readings; try to avoid producing air bubbles and always remove bubbles prior to reading plates |
|  | Concentration of standard stock incorrect | Recheck datasheet for recommended concentrations of standard stocks |
|  | Errors in standard curve calculations | Refer to datasheet and re-check the calculations |
|  | Use of other reagents than those provided with the kit | Use fresh components from the same kit |

For further technical questions please do not hesitate to contact us by email (technical@abcam.com) or phone (select “contact us” on www.abcam.com for the phone number for your region).

Technical Support

Copyright © 2023 Abcam. All Rights Reserved. The Abcam logo is a registered trademark. All information / detail is correct at time of going to print.

For all technical or commercial enquiries please go to:

www.abcam.com/contactus

www.abcam.cn/contactus (China)

www.abcam.co.jp/contactus (Japan)