ab209883 – Mouse Angiopoietin 2 (ANG2) SimpleStep ELISA® Kit

Instructions for use:

For the quantitative measurement of Angiopoietin 2 (ANG2) in mouse serum, plasma and cell culture supernatants samples.

This product is for research use only and is not intended for diagnostic use.

Table of Contents

INTRODUCTION	1

1.	BACKGROUND	1

2.	ASSAY SUMMARY	2

GENERAL INFORMATION	3

3.	PRECAUTIONS	3

4.	STORAGE AND STABILITY	3

5.	MATERIALS SUPPLIED	3

6.	MATERIALS REQUIRED, NOT SUPPLIED	4

7.	LIMITATIONS	4

8.	TECHNICAL HINTS	4

ASSAY PREPARATION	6

9.	REAGENT PREPARATION	6

10.	STANDARD PREPARATION	7

11.	SAMPLE PREPARATION	8

12.	PLATE PREPARATION	9

ASSAY PROCEDURE	10

13.	ASSAY PROCEDURE	10

DATA ANALYSIS	12

14.	CALCULATIONS	12

15.	TYPICAL DATA	13

16.	TYPICAL SAMPLE VALUES	14

17.	ASSAY SPECIFICITY	18

18.	SPECIES REACTIVITY	19

RESOURCES	20

19.	TROUBLESHOOTING	20

20.	NOTES	21

INTRODUCTION

# BACKGROUND

Abcam’s Angiopoietin 2 (ANG2) in vitro SimpleStep ELISA® (Enzyme-Linked Immunosorbent Assay) kit is designed for the quantitative measurement of Angiopoietin 2 protein in mouse serum, plasma and cell culture supernatants samples.

The SimpleStep ELISA® employs an affinity tag labeled capture antibody and a reporter conjugated detector antibody which immunocapture the sample analyte in solution. This entire complex (capture antibody/analyte/detector antibody) is in turn immobilized via immunoaffinity of an anti-tag antibody coating the well. To perform the assay, samples or standards are added to the wells, followed by the antibody mix. After incubation, the wells are washed to remove unbound material.  TMB substrate is added and during incubation is catalyzed by HRP, generating blue coloration. This reaction is then stopped by addition of Stop Solution completing any color change from blue to yellow. Signal is generated proportionally to the amount of bound analyte and the intensity is measured at 450 nm. Optionally, instead of the endpoint reading, development of TMB can be recorded kinetically at 600 nm.

Angiopoietin 2 is a secreted glycosylated protein functioning in angiogenesis, inflammation and vascular development. Angiopoietin 2 binds to receptor tyrosine kinase TEK/TIE2, competing for the ANGPT1 binding site, and modulating ANGPT1 signalling. Angiopoietin 2 can induce tyrosine phosphorylation of TEK/TIE2 in the absence of ANGPT1. In the absence of angiogenic inducers, such as VEGF, Angiopoietin 2-mediated loosening of cell-matrix contacts may induce endothelial cell apoptosis with consequent vascular regression. In concert with VEGF, Angiopoietin 2 may facilitate endothelial cell migration and proliferation, thus serving as a permissive angiogenic signal.

# ASSAY SUMMARY

Remove appropriate number of antibody coated well strips. Equilibrate all reagents to room temperature.  Prepare all reagents, samples, and standards as instructed.

Add standard or sample to appropriate wells.

Add Antibody Cocktail to all wells. Incubate at room temperature.

Aspirate and wash each well. Add TMB Substrate to each well and incubate. Add Stop Solution at a defined endpoint. 
Alternatively, record color development kinetically after TMB substrate addition.

GENERAL INFORMATION

# PRECAUTIONS

Please read these instructions carefully prior to beginning the assay.

All kit components have been formulated and quality control tested to function successfully as a kit.  Modifications to the kit components or procedures may result in loss of performance.

# STORAGE AND STABILITY

Store kit at 2-8ºC immediately upon receipt.

Refer to list of materials supplied for storage conditions of individual components.  Observe the storage conditions for individual prepared components in the Reagent and Standard Preparation sections.

# MATERIALS SUPPLIED

| Item | Amount | Storage Condition (Before Preparation) |
| --- | --- | --- |
| 10X Mouse Angiopoietin 2 Capture Antibody | 600 µL | +2-8ºC |
| 10X Mouse Angiopoietin 2 Detector Antibody | 600 µL | +2-8ºC |
| Mouse Angiopoietin 2 Lyophilized Recombinant Protein | 2 Vials | +2-8ºC |
| Antibody Diluent 4BR | 6 mL | +2-8ºC |
| 10X Wash Buffer PT | 20 mL | +2-8ºC |
| TMB Substrate | 12 mL | +2-8ºC |
| Stop Solution | 12 mL | +2-8ºC |
| Sample Diluent NS | 50 mL | +2-8ºC |
| Pre-Coated 96 Well Microplate (12 x 8 well strips) | 96 Wells | +2-8ºC |
| Plate Seal | 1 | +2-8ºC |

# MATERIALS REQUIRED, NOT SUPPLIED

These materials are not included in the kit, but will be required to successfully utilize this assay:

- Microplate reader capable of measuring absorbance at 450 or 600 nm.

- Method for determining protein concentration (BCA assay recommended).

- Deionized water.

- Multi- and single-channel pipettes.

- Tubes for standard dilution.

- Plate shaker for all incubation steps.

- Optional: Phenylmethylsulfonyl Fluoride (PMSF) (or other protease inhibitors).

# LIMITATIONS

Assay kit intended for research use only. Not for use in diagnostic procedures.

Do not mix or substitute reagents or materials from other kit lots or vendors. Kits are QC tested as a set of components and performance cannot be guaranteed if utilized separately or substituted.

# TECHNICAL HINTS

- Samples generating values higher than the highest standard should be further diluted in the appropriate sample dilution buffers.

- Avoid foaming or bubbles when mixing or reconstituting components.

- Avoid cross contamination of samples or reagents by changing tips between sample, standard and reagent additions.

- Ensure plates are properly sealed or covered during incubation steps.

- Complete removal of all solutions and buffers during wash steps is necessary to minimize background.

- As a guide, typical ranges of sample concentration for commonly used sample types are shown below in Sample Preparation (section 11).

- All samples should be mixed thoroughly and gently.

- Avoid multiple freeze/thaw of samples.

- Incubate ELISA plates on a plate shaker during all incubation steps.

- When generating positive control samples, it is advisable to change pipette tips after each step.

- To avoid high background always add samples or standards to the well before the addition of the antibody cocktail.

- This kit is sold based on number of tests.  A ‘test’ simply refers to a single assay well.  The number of wells that contain sample, control or standard will vary by product.  Review the protocol completely to confirm this kit meets your requirements.  Please contact our Technical Support staff with any questions.

ASSAY PREPARATION

# REAGENT PREPARATION

- Equilibrate all reagents to room temperature (18-25°C) prior to use.  The kit contains enough reagents for 96 wells.  The sample volumes below are sufficient for 48 wells (6 x 8-well strips); adjust volumes as needed for the number of strips in your experiment.

- Prepare only as much reagent as is needed on the day of the experiment.  Capture and Detector Antibodies have only been tested for stability in the provided 10X formulations.

  - 1X Wash Buffer PT

Prepare 1X Wash Buffer PT by diluting 10X Wash Buffer PT with deionized water.  To make 50 mL 1X Wash Buffer PT combine 5 mL 10X Wash Buffer PT with 45 mL deionized water.  Mix thoroughly and gently.

  - Antibody Cocktail

Prepare Antibody Cocktail by diluting the capture and detector antibodies in Antibody Diluent 4BR.  To make 3 mL of the Antibody Cocktail combine 300 µL 10X Capture Antibody and 300 µL 10X Detector Antibody with 2.4 mL Antibody Diluent 4BR.  Mix thoroughly and gently.

# STANDARD PREPARATION

Prepare serially diluted standards immediately prior to use.  Always prepare a fresh set of positive controls for every use.

The following section describes the preparation of a standard curve for duplicate measurements (recommended).

IMPORTANT: If the protein standard vial has a volume identified on the label, reconstitute the Angiopoietin 2 standard by adding that volume of Sample Diluent NS indicated on the label. Alternatively, if the vial has a mass identified, reconstitute the Angiopoietin 2 standard by adding 500 µL Sample Diluent NS. Hold at room temperature for 10 minutes and mix gently.  This is the 10,560 pg/mL Stock Standard Solution.

Label eight tubes, Standards 1– 8.

Add 175 μL Sample Diluent NS into tube number 1 and 150 μL of Sample Diluent NS into numbers 2-8.

Use the Stock Standard to prepare the following dilution series. Standard #8 contains no protein and is the Blank control:

10,560

pg/mL

4,400

pg/mL

2,200

pg/mL

1,100

pg/mL

550

pg/mL

275

pg/mL

137.5

pg/mL

68.75

pg/mL

125 µL

150 µL

µ

150 µL

µ

150 µL

µ

150 µL

µ

150 µL

µ

150 µL

µ

0

pg/mL

# SAMPLE PREPARATION

| TYPICAL SAMPLE DYNAMIC RANGE |  |
| --- | --- |
| Sample Type | Range |
| Mouse Plasma - Heparin | 1:107 – 1:13 |
| Mouse Plasma - EDTA | 1:107 – 1:13 |
| Mouse Plasma - Citrate | 1:267 – 1:17 |
| Mouse Serum | 1:32 – 1:4 |
| Rat Serum | 1:32 – 1:4 |

Plasma
Collect plasma using citrate, EDTA or heparin. Centrifuge samples at 2,000 x g for 10 minutes. Dilute plasma (heparin, EDTA) samples at least 1:13 into Sample Diluent NS and assay.  Dilute plasma (citrate) samples at least 1:17 into Sample Diluent NS and assay.  Store un-diluted plasma samples at -20ºC or below for up to 3 months. Avoid repeated freeze-thaw cycles.

Serum
Samples should be collected into a serum separator tube. After clot formation, centrifuge samples at 2,000 x g for 10 minutes and collect serum. Dilute serum samples at least 1:4 into Sample Diluent NS and assay.   Store un-diluted serum at -20ºC or below. Avoid repeated freeze-thaw cycles.

  - Cell Culture Supernatants

Centrifuge cell culture media at 2,000 x g for 10 minutes to remove debris. Collect supernatants. Supernatant samples can be assayed without dilution. If needed, dilute supernatant samples into Sample Diluent NS and assay. Store un-diluted samples at -20°C or below. Avoid repeated freeze-thaw cycles.

# PLATE PREPARATION

The 96 well plate strips included with this kit are supplied ready to use.  It is not necessary to rinse the plate prior to adding reagents.

Unused plate strips should be immediately returned to the foil pouch containing the desiccant pack, resealed and stored at 4°C.

For each assay performed, a minimum of two wells must be used as the zero control.

For statistical reasons, we recommend each sample should be assayed with a minimum of two replicates (duplicates).

Differences in well absorbance or “edge effects” have not been observed with this assay.

ASSAY PROCEDURE

# ASSAY PROCEDURE

- Equilibrate all materials and prepared reagents to room temperature prior to use.

- It is recommended to assay all standards, controls and samples in duplicate.

Prepare all reagents, working standards, and samples as directed in the previous sections.

Remove excess microplate strips from the plate frame, return them to the foil pouch containing the desiccant pack, reseal and return to 4ºC storage.

Add 50 µL of all sample or standard to appropriate wells.

Add 50 µL of the Antibody Cocktail to each well.

Seal the plate and incubate for 1 hour at room temperature on a plate shaker set to 400 rpm.

Wash each well with 3 x 350 µL 1X Wash Buffer PT.  Wash by aspirating or decanting from wells then dispensing 350 µL 1X Wash Buffer PT into each well.  Complete removal of liquid at each step is essential for good performance.  After the last wash invert the plate and blot it against clean paper towels to remove excess liquid.

Add 100 µL of TMB Development Solution to each well and incubate for 10 minutes in the dark on a plate shaker set to 400 rpm.

- Given variability in laboratory environmental conditions, optimal incubation time may vary between 5 and 20 minutes.
Note: The addition of Stop Solution will change the color from blue to yellow and enhance the signal intensity about 3X. To avoid signal saturation, proceed to the next step before the high concentration of the standard reaches a blue color of O.D.600 equal to 1.0.

Add 100 µL of Stop Solution to each well.  Shake plate on a plate shaker for 1 minute to mix.  Record the OD at 450 nm.  This is an endpoint reading.

Alternative to 13.7 – 13.8:  Instead of the endpoint reading at 450 nm, record the development of TMB Substrate kinetically.  Immediately after addition of TMB Development Solution begin recording the blue color development with elapsed time in the microplate reader prepared with the following settings:

| Mode: | Kinetic |
| --- | --- |
| Wavelength: | 600 nm |
| Time: | up to 20 min |
| Interval: | 20 sec - 1 min |
| Shaking: | Shake between readings |

Note that an endpoint reading can also be recorded at the completion of the kinetic read by adding 100 µL Stop Solution to each well and recording the OD at 450 nm.

Analyze the data as described below.

DATA ANALYSIS

# CALCULATIONS

  - Calculate the average absorbance value for the blank control (zero) standards. Subtract the average blank control standard absorbance value from all other absorbance values.

  - Create a standard curve by plotting the average blank control subtracted absorbance value for each standard concentration (y-axis) against the target protein concentration (x-axis) of the standard. Use graphing software to draw the best smooth curve through these points to construct the standard curve.

Note: Most microplate reader software or graphing software will plot these values and fit a curve to the data.  A four parameter curve fit (4PL) is often the best choice; however, other algorithms (e.g. linear, semi-log, log/log, 4 parameter logistic) can also be tested to determine if it provides a better curve fit to the standard values.

  - Determine the concentration of the target protein in the sample by interpolating the blank control subtracted absorbance values against the standard curve.  Multiply the resulting value by the appropriate sample dilution factor, if used, to obtain the concentration of target protein in the sample.

  - Samples generating absorbance values greater than that of the highest standard should be further diluted and reanalyzed.  Similarly, samples which measure at an absorbance values less than that of the lowest standard should be retested in a less dilute form.

# TYPICAL DATA

TYPICAL STANDARD CURVE – Data provided for demonstration purposes only.  A new standard curve must be generated for each assay performed.

| Standard Curve Measurements |  |  |  |
| --- | --- | --- | --- |
| Conc. | O.D. 450 nm |  | Mean |
| (pg/mL) | 1 | 2 | O.D. |
| 0 | 0.145 | 0.117 | 0.131 |
| 68.75 | 0.181 | 0.186 | 0.183 |
| 137.5 | 0.261 | 0.251 | 0.256 |
| 275 | 0.341 | 0.380 | 0.360 |
| 550 | 0.665 | 0.577 | 0.621 |
| 1,100 | 0.980 | 1.066 | 1.023 |
| 2,200 | 1.794 | 1.959 | 1.876 |
| 4,400 | 3.710 | 3.523 | 3.616 |

Figure 1. Example of mouse Angiopoietin 2 standard curve.  The Angiopoietin 2 standard curve was prepared as described in Section 10.  Raw data values are shown in the table.  Background-subtracted data values (mean +/- SD) are graphed.

# TYPICAL SAMPLE VALUES

SENSITIVITY –

The calculated minimal detectable dose (MDD) is 17.3 pg/mL.  The MDD was determined by calculating the mean of zero standard replicates (n=40) and adding 2 standard deviations then extrapolating the corresponding concentration.

RECOVERY –

Three concentrations of mouse recombinant Angiopoietin 2 were spiked in duplicate to the indicated biological matrix to evaluate signal recovery in the working range of the assay.

| Sample Type | Average % Recovery | Range (%) |
| --- | --- | --- |
| 10% Mouse Serum | 94.5 | 67.4 – 120.9 |
| 6% Mouse Plasma - Citrate | 94.7 | 87.1 – 103.3 |
| 6% Mouse Plasma - EDTA | 104.9 | 91.8 – 122.8 |
| 6% Mouse Plasma - Heparin | 112.4 | 81.4 – 143.2 |
| 100% Cell Culture Media Containing 10% Bovine Serum | 99.7 | 94.3 – 103.3 |

LINEARITY OF DILUTION –

Linearity of dilution is determined based on interpolated values from the standard curve.  Linearity of dilution defines a sample concentration interval in which interpolated target concentrations are directly proportional to sample dilution.

Native Angiopoietin 2 was measured in the following biological samples in a 2-fold dilution series.  Sample dilutions are made in Sample Diluent NS.

| Dilution Factor | Interpolated value | 25% Mouse Serum | 25% Rat Serum | 6% Mouse Plasma (Citrate) | 7.5% Mouse Plasma (EDTA) | 7.5% Mouse Plasma (Heparin) |
| --- | --- | --- | --- | --- | --- | --- |
| Undiluted | pg/mL | 4,654.6 | 695.7 | 1,934.1 | 2,077.0 | 2,552 |
|  | % Expected value | 100 | 100 | 100 | 100 | 100 |
| 2 | pg/mL | 2,767.6 | 429 | 985.3 | 961.3 | 1,270.8 |
|  | % Expected value | 118.9 | 123.3 | 101.9 | 92.6 | 99.6 |
| 4 | pg/mL | 1,203.7 | 194.7 | 493.8 | 538 | 711.8 |
|  | % Expected value | 103.4 | 112.0 | 102.1 | 103.6 | 111.6 |
| 8 | pg/mL | 501.6 | 68.8 | 241.4 | 222.8 | 276.9 |
|  | % Expected value | 86.2 | 79.1 | 99.8 | 85.8 | 86.8 |
| 16 | pg/mL | NL | NA | 96.9 | NL | NL |
|  | % Expected value | NL | NA | 80.2 | NL | NL |

NL – Non-Linear; NA – Not Analyzed (O.D. values lower than the O.D. value of the lowest standard

Recombinant mouse Angiopoietin 2 was spiked into the following biological samples and diluted in a 2-fold dilution series in Sample Diluent NS.

| Dilution Factor | Interpolated value | 100% Cell Culture Media |
| --- | --- | --- |
| Undiluted | pg/mL | 3,994.6 |
|  | % Expected value | 100 |
| 2 | pg/mL | 2,090.9 |
|  | % Expected value | 104.7 |
| 4 | pg/mL | 992.2 |
|  | % Expected value | 99.4 |
| 8 | pg/mL | 488 |
|  | % Expected value | 97.7 |
| 16 | pg/mL | 274.8 |
|  | % Expected value | 110.1 |

PRECISION –

Mean coefficient of variations of interpolated values of Angiopoietin 2 in 3 concentrations of mouse plasma (citrate) within the working range of the assay.

|  | Intra- Assay | Inter- Assay |
| --- | --- | --- |
| n= | 3 | 5 |
| CV (%) | 8.3 | 10.3 |

Figure 2. Interpolated concentrations of native Angiopoietin 2 in mouse and rat serum and mouse plasma samples.  The concentrations of Angiopoietin 2 were measured in duplicates, interpolated from the Angiopoietin 2 standard curves and corrected for sample dilution. Undiluted samples are as follows: serum 25%, plasma (citrate) 6%, plasma (EDTA) 7.5%, plasma (heparin) 7.5%.  The interpolated dilution factor corrected values are plotted (mean +/- SD, n=2).  The mean Angiopoietin 2 concentration was determined to be 19,017 pg/mL in mouse serum, 2,802 pg/mL in rat serum, 31,192 pg/mL in mouse plasma (citrate), 26,451 pg/mL in mouse plasma (EDTA) and 33,857 pg/mL in mouse plasma (heparin).

# ASSAY SPECIFICITY

This kit recognizes both native and recombinant mouse/rat Angiopoietin 2 protein in serum, plasma and cell culture supernatant samples only.

Tissue extract samples are not compatible with this kit.

Urine, milk, saliva and cell extract samples have not been tested with this kit.

CROSS REACTIVITY

Recombinant human Angiopoietin 1 was prepared at 10,200 pg/mL and assayed for cross reactivity.  No cross-reactivity was observed.

INTERFERENCE

Recombinant human Angiopoietin 1 at 9,200 pg/mL was tested for interference with 7-point 2-fold serial dilution of mouse Angiopoietin 2 starting at 4,600 pg/mL.  No interference with was observed.

# SPECIES REACTIVITY

This kit recognizes mouse and rat Angiopoietin 2 protein.

Recombinant human Angiopoietin 2 was prepared at 5,100 pg/mL, 2,550 pg/mL and 1,275 pg/mL and assayed for reactivity. Only 4.1% reactivity was observed.

Human and cow species reactivity was determined by measuring 25% serum samples of these species, interpolating the protein concentrations from the mouse/rat standard curve, and expressing the interpolated concentrations as a percentage of the protein concentration in mouse serum assayed at the same dilution.

Reactivity < 3% was determined for the following species:

- Human

- Cow

Please contact our Technical Support team for more information.

RESOURCES

# TROUBLESHOOTING

| Problem | Cause | Solution |
| --- | --- | --- |
| Poor standard curve | Inaccurate Pipetting | Check pipettes |
|  | Improper standard dilution | Prior to opening, briefly spin the stock standard tube and dissolve the powder thoroughly by gentle mixing |
| Low Signal | Incubation times too brief | Ensure sufficient incubation times; increase to 2 or 3 hour standard/sample incubation |
|  | Inadequate reagent volumes or improper dilution | Check pipettes and ensure correct preparation |
|  | Incubation times with TMB too brief | Ensure sufficient incubation time until blue color develops prior addition of Stop solution |
| Large CV | Plate is insufficiently washed | Review manual for proper wash technique. If using a plate washer, check all ports for obstructions. |
|  | Contaminated wash buffer | Prepare fresh wash buffer |
| Low sensitivity | Improper storage of the ELISA kit | Store your reconstituted standards at -80°C, all other assay components 4°C. Keep TMB substrate solution protected from light. |
| Precipitate in Diluent | Precipitation and/or coagulation of components within the Diluent. | Precipitate can be removed by gently warming the Diluent to 37ºC. |

# NOTES

Technical Support

Copyright © 2022 Abcam. All Rights Reserved. The Abcam logo is a registered trademark. All information / detail is correct at time of going to print.

For all technical or commercial enquiries please go to:

www.abcam.com/contactus

www.abcam.cn/contactus (China)

www.abcam.co.jp/contactus (Japan)