Version 3a, Last updated 6 June 2023

# ab235696 # Cell Migration/Chemotaxis Assay Kit (24-well, 5 µm)

For the measurement of cell migration in response to stimuli.

This product is for research use only and is not intended for diagnostic use.

PLEASE NOTE: With the acquisition of BioVision by Abcam, we have made some changes to component names and packaging to better align with our global standards as we work towards environmental-friendly and efficient growth. You are receiving the same high-quality products as always, with no changes to specifications or protocols.

Table of Contents

1.	Overview	1

2.	Materials Supplied and Storage	2

3.	Materials Required, Not Supplied	3

4.	General guidelines, precautions, and troubleshooting	4

5.	Reagent Preparation	5

6.	Assay Procedure	6

7.	Data Analysis	10

8.	FAQs / Troubleshooting	11

9.	Typical Data	12

10.	Notes	14

Overview

Cell Migration/Chemotaxis Assay Kit (24-well, 5 µm) (ab235696) utilizes a Boyden chamber, where the cells migrate through a semi-permeable membrane under different stimuli. Cell migration can be analyzed directly by reading fluorescence (Ex/Em = 530/590 nm) in a plate reader. Our assay is easy to use, sensitive and adaptable to high-throughput systems.

Prepare cells.

Prior to the assay, starve cells for 18-24 h in serum-free media.

Set up cell migration assay containing desired chemoattractant in the bottom chamber. Incubate the Cell Migration Chamber (12 x 5µm)/Cell Migration Chamber at 37°C in CO2 incubator for 2-48 h.

Prepare Standard Curve for each cell type.

Separate Migrated cells.

Add cell dye and count Migrated cells.

Materials Supplied and Storage

Store kit at 20°C in the dark immediately on receipt and check below for storage for individual components. Kit can be stored for 1 year from receipt, if components have not been reconstituted.

Components are stable for 6 months after the first thaw.

Aliquot components in working volumes before storing at the recommended temperature.

Avoid repeated freeze-thaws of reagents.

| Item | Quantity | Storage temperature (before prep) | Storage temperature (after prep) |
| --- | --- | --- | --- |
| Wash Buffer I/Wash Buffer | 25 mL | -20°C | -20°C |
| Cell Dissociation Solution II/Cell Dissociation Solution | 6 mL | -20°C | -20°C |
| Control Migration Inducer | 300 µL | -20°C | -20°C |
| Cell Dye I/Cell Dye | 1.5 mL | -20°C | -20°C |
| Cell Migration Chamber (12 x 5µm)/Cell Migration Chamber | 1 | -20°C | -20°C |

Materials Required, Not Supplied

These materials are not included in the kit, but will be required to successfully perform this assay:

- Fluorescence Plate Reader.

- Cotton Swab.

- Centrifuge to spin 96-well plate.

- 96-Well White Plates with clear bottom.

General guidelines, precautions, and troubleshooting

Please observe safe laboratory practice and consult the safety datasheet.

For general guidelines, precautions, limitations on the use of our assay kits and general assay troubleshooting tips, particularly for first time users, please consult our guide: www.abcam.com/assaykitguidelines

For typical data produced using the assay, please see the assay kit  datasheet on our website.

Reagent Preparation

Briefly centrifuge small vials at low speed prior to opening.

  - Wash Buffer I/Wash Buffer

- Ready to use as supplied.

- Bring to 37°C before use.

- Stable for six months after the first thaw.

  - Cell Dissociation Solution II/Cell Dissociation Solution

- Ready to use as supplied.

- Bring to 37°C before use.

- Stable for six months after the first thaw.

  - Control Migration Inducer

- Ready to use as supplied.

- Bring to 37°C before use.

- Stable for six months after the first thaw.

  - Cell Dye I/Cell Dye

- Ready to use as supplied.

- Aliquot and store at -20°C.

- Bring to 37°C before use.

  - Cell Migration Chamber (12 x 5µm)/Cell Migration Chamber

- Open under sterile conditions.

Assay Procedure

- Equilibrate all materials and prepared reagents to room temperature just prior to use and gently agitate.

- Assay all standards, controls and samples in duplicate.

  - Cell Migration Assay Protocol:

- Grow enough cells to perform a Cell Migration Assay and a Standard Curve in desired media and culture conditions.

- Adherent cells should be cultured to ~80% confluence. Prior to the assay, starve cells for 18-24 h  in serum-free media (0.5% serum can be used if needed).

- After starvation, harvest cells and centrifuge at 1,000 x g, for 5 minutes to pellet them.

- Resuspend the cell pellet in serum-free media and count the number of cells using a hemocytometer or an automated cell counter.

- Resuspend cells at 1 x 106 cells/mL in a serum-free media.

- Under sterile conditions, disassemble the Cell Migration Chamber (12 x 5µm)/Cell Migration Chamber (Figure 1) and carefully remove the plate cover and the top chamber.

Figure 1. Cell Invasion plate: The cells are added to the Top Chamber and the Control Invasion Inducer or chemoattractant are added to the Bottom Chamber.

- Bottom Chamber: Add 600 µL of serum-free media per well containing desired chemoattractant to the bottom chamber.

- In control well(s), we recommend omitting the chemoattractant.

- For positive control, add 60 µL of Control Migration Inducer to 540 µL of media in the bottom chamber.

- Place the top chamber back into the bottom chamber.

- Ensure no air bubbles are trapped between the top and the bottom chamber.

- Top Chamber: Add 200 µL (2-3 x 105 cells) of cell suspension to each well of the top chamber.

- Add desired stimulator or inhibitor to the top well, and gently mix.

- Carefully place the plate cover and incubate the Cell Migration Chamber (12 x 5µm)/Cell Migration Chamber at 37°C in CO2 incubator for 2-48 hours.

Note: Migratory cells pass through the polycarbonate membrane and/or cling to the outer side of the top chamber. Non-migratory cells stay in the upper chamber.

Note: If required, media with 0.1% serum can be used in top chamber.

  - Standard Curve:

- Each cell type requires a separate Standard Curve. Prepare a Standard Curve by adding 50 µL cell suspension (1 x 106 cells/mL, 50,000 cells) per well in 96-well white plates with a clear bottom.

- Serially dilute the cells 1:1 in Wash Buffer I/Wash Buffer and generate a Standard Curve of cells (50,000, 25,000, 12,500, 6,250, 3,125, 1,562, 781 and 390) in 100 µL total volume.

- As blank, use 100 µL of Wash Buffer I/Wash Buffer.

- Add 10 µL of Cell Dye I/Cell Dye to each well.

- Incubate at 37°C for 1 hour.

- Read the fluorescence at Ex/Em = 530/590 nm.

- Plot the Standard Curve of Number of Cells Vs RFU obtained.

- Fit the data points using a linear trendline with zero intercept.

- The equation for the straight line and R-square value are used for data analysis of samples.

Note: The Cell Migration RFU reading should fall in the linear range of the Standard Curve. We recommend using triplicates for Standard Curve.

  - Separation of Invasive Cells:

- After the desired incubation with cell invasion inducers/inhibitors, carefully remove the plate cover and aspirate media from the top chamber without puncturing the membrane and matrix.

- Remove cells from the top chamber using a cotton swab. Disassemble the Cell Invasion Chamber by removing the top chamber. Invert the top chamber and set it aside.

- Place the plate cover on top of bottom chamber and centrifuge the plate at 1,000 x g for 5 minutes at room temperature.

- Carefully aspirate the media from the bottom chamber, and wash the chamber with 500 µL Wash Buffer I/Wash Buffer.

- Centrifuge the plate at 1,000 x g for 5 min. at room temperature and aspirate the Wash Buffer I/Wash Buffer from the bottom chamber.

  - Count Invasive Cells:

- For every twenty wells to be assayed, prepare a mix of 1 mL of Cell Dye I/Cell Dye in 10 mL of Cell Dissociation Solution II/Cell Dissociation Solution. Mix well. Make the Cell Dye I/Cell Dye solution as desired depending on the number of wells.

- Add 550 µL of the mix to each well of the bottom chamber. Reassemble the Cell Invasion Chamber by placing the top chamber into the bottom chamber. Incubate at 37°C in CO2 incubator for 60 minutes.

- After incubation, disassemble the Cell Invasion Chamber, remove the top chamber and transfer 110 µL of mix from the bottom chamber to the 96-well white plate (the same plate having Standards).

- Read the plate at Ex/Em = 530/590 nm. Multiply the reading by 5 to account for the 5X higher volume in each well of the 24-well plate.

Data Analysis

- Calculate the number of cells invaded using the equation of the straight line obtained from Standard Curve.

- Percentage Invasion can be calculated as follows:

% Invasion=B# Cells in Lower Chamber  Total # Cells added to Top Chamber  *100

FAQs / Troubleshooting

General troubleshooting points are found at www.abcam.com/assaykitguidelines

Typical Data

Data provided for demonstration purposes only.

Figure 2. Standard Curve: Monocytes/macrophage cells were harvested, counted and serially diluted to obtain desired cell number. Cells were incubated according to the protocol.

Figure 3. Cell Invasion: Monocytes/macrophage cells were starved overnight and treated with Control (Cnt) Invasion Inducer for 24 or 48 hours or left untreated. Treatment with Control Invasion Inducer demonstrated a significant increase in invasion with time. Control reading was subtracted from inducer reading.

Notes

# Technical Support

Copyright © 2023 Abcam. All Rights Reserved. The Abcam logo is a registered trademark. All information / detail is correct at time of going to print.

For all technical or commercial enquiries please go to:

www.abcam.com/contactus

www.abcam.cn/contactus (China)

www.abcam.co.jp/contactus (Japan)